OSHA ID-110 Fluoride (F- and HF) in Workplace Atmospheres

- Summary of Method Modifications:

The method being modified is OSHA ID-110 for Fluoride analysis. Updates for equipment used, and sample preparation clarification and modifications were necessary to improve efficiency of the method. These modifications are supported by data collected during testing and are found in Appendix A of this Method Work Instruction.

- Equipment Modifications:

- A benchtop Thermo Scientific Orion Dual Star pH/ISE meter (Cat. #13-642-780) is currently used instead of an Orion Research Inc. EA 940 Expandable Ionizer.

- A Fisher Scientific Accumet pencil-thin epoxy body gel-filled combination electrode (Cat. #13-620-290) is used for pH adjustments, no filling solution is needed.

- A Thermo Scientific Orion ion-plus sure-flow Fluoride electrode (Cat. #9609BNWP) is used. No reference electrodes are used for F- analysis.

- Hamilton Auto Diluter is used for standards preparation. Hamilton Auto Diluter or dispenser bottle (calibrate desired volumes prior to use) can be used for Tris-Tartrate buffer addition.

- SCP Class A 50mL DigiTubes are used for samples and standards preparation.

- Eppendorf pipettes are used during HF sample desorption.

- Glassware and plasticware are no longer rinsed or acid washed before use.

- Polyethylene beakers are no longer used for fluoride analysis.

- Reagents and Standards Modifications:

- Tris(hydroxymethyl)aminomethane CAS #77-86-1 is used to prepare the Tris-Tartrate Buffer (concentrated).

- The Tris-Tartrate Buffer (concentrated) can be sonicated for the solute to properly dissolve. The mixture will produce an endothermic reaction and must be allowed to reach room temperature before bringing it up to volume in the volumetric flask.

- Working standards and ICV are prepared from 100 ppm intermediate F- standards. Intermediate F- standards 100 ppm are prepared from 10X dilutions of 1000 ppm F- stock standards. Dilutions are done using a Hamilton Auto Diluter with dilution programs for 5X, 10X, 50X, and 100X for working standards and 10X for ICV.

- Analytical Parameter Modifications:

Sample Preparation:

HF (Na2CO3 coated backup pads):

- Place the backup pad into a scintillation vial and add approximately 20 mL of 18 MΩ DI H2O. Allow the samples to sit for an hour, shaking them twice during this time. Transfer the sample solution to a SCP Class A 50 mL DigiTube. Adjust the pH of the sample solution to between 4.5-10.5 using 2% HCl and dilute them to 25 mL with 18 MΩ DI H2O. This method modification resulted in percent recoveries within established quality control limits at reporting limit levels and QCSM spiked levels as shown in Appendix A.

Note: The MCEF should always be prepared and analyzed even if the industrial hygienist specifies only an HF analysis. A set of QCSMs is analyzed alongside with samples (HF and particulate F-).

Particulate F- (MCEF):

- Place the sample filter in a 75 mL nickel crucible. Using tape, secure the crucibles to a metal rack, separating them at least two inches from each other to prevent cross-contamination from possible splattering during heating.

- Add 5 mL of 5N NaOH using an Eppendorf pipette.

- Turn on the vacuum-assisted drying oven at 90 ºC. Once it reaches temperature, place the samples in the bottom rack of the oven, and adjust the temperature to 140 ºC. The temperature will slowly ramp up; allow samples to dry for an hour at 140 ºC.

- After an hour, open the oven vents. Then, turn on the vacuum to the drying oven and close the vent. Allow samples to dry overnight.

- After drying has finished, open the vent on the oven, turn off the vacuum and the oven. When the oven has reached ambient pressure, the door can be opened. Remove the rack with the samples. The crucibles will be hot and should be allowed to cool until they can be handled to remove any tape used for securing prior to fusing.

- After the samples have been fused as described in the method, allow them to cool, add approximately 10 mL of 18 MΩ DI H2O, and add 2 mL of concentrated HCl. Swirl the samples until all the white deposit on the bottom of the crucible has dissolved.

- Transfer the sample solution to a SCP Class A 50 mL DigiTube, adjust them for pH with concentrated HCl, and dilute them up to the 25 mL mark with DI water on the DigiTube. If the volume is inadvertently taken past the 25 mL mark, dilute to the next graduation on the Digitube.

Sample aliquots can be taken in case the analyst wishes to reserve sample solution for a rerun or if volumes are taken past the final volume of 25 mL. Document the aliquot volume so that the same amount of Tris buffer can be added analytically just before analysis. The aliquot volume is needed to calculate the dilution factor to be used for the sample result.

Analysis:

- The addition of 1:1 Tris buffer (analytical) to standards and samples is performed right before the analytical run.

- A standard at the reporting limit is analyzed after initial calibration and at the end of the analysis.

- The analysis is done strictly by using standard addition program in the ISE meter.

- Instrument Set-up Modifications and Instructions:

Electrode preparation:

- The ISE/pH meter instrument will require electrode stabilization time. Use the power button to turn the meter on. Channel 1 is set up for fluoride, while Channel 2 is set up for pH. For pH, submerge the electrode in a buffer solution and allow to stabilize for at least an hour. For fluoride, submerge the electrode in a 1 ppm standard solution and allow to stabilize for at least an hour. After stabilizing, electrodes must be calibrated as described next.

pH Calibration:

- Press the “Cal” button and select Channel 2 using the arrows. The number “2” button is also the “up” arrow, while the number “3” button is the “down” arrow. Follow the prompts on the screen, they will guide the analyst to rinse the electrode, place it in buffer solution and press start.

- pH 4.01 and 10.01 are used for pH calibration. The ISE will start reading the pH of the buffer solution. Once stable, a message will appear stating “Accept or enter value”. Enter the buffer’s pH value from the manufacturer and press accept. Press “next” and repeat the previous steps with the next buffer solution. When finished, press “cal. done”, and record the slope value (this is to trace instrument performance). Press “log/print” for the ISE to store the calibration data for the run.

ISE Calibration:

- ISE calibration is performed with neat standards without addition of 1:1 Tris buffer (analytical) solution. Standard 1 and standard 10 are used for ISE calibration.

- Press the “Cal” button and select Channel 1 using the arrows. The number “2” button is also the “up” arrow, while the number “3” button is the “down” arrow. Follow the prompts on the screen, they will guide the analyst to rinse the electrode, place it in the standard solution (Std. 1) and press start.

- The ISE will start reading the F- concentration in the standard. Once stable, a message will appear stating “Accept or enter value”. Enter the standard theoretical value (typing in the theoretical value will override the measured value displayed) and press accept.

- Press “next” and repeat the previous steps with the next standard (Std. 10). When, finished press “cal. done”, and record the slope value.

- The value must be between - 54 to - 60 mV/dec. This slope value will be used for the standard addition program of the analytical run. Press “log/print” for the ISE to store the calibration data.

ISE Standard Addition Program:

- Before the start of the run, add 1:1 amounts of Tris Buffer (Analytical) to each standard and sample.

- On the ISE instrument, press the “Set Up” key. Then, select “Incremental Techniques”, “Channel 1”, and “Single Known Addition”. This will initiate the program for standard addition in the ISE.

- First, enter the value for the slope acquired during the Fluoride electrode calibration, press “accept” after each entry to go to the next step. Next, enter the total sample volume in mL for both the “sample volume” and “total volume” screens; this volume will usually be 40 mL for standards, and 50 mL for QCs and samples.

- Then, enter the “standard concentration” from which standard addition will be performed; this concentration value will usually be 1000 mg/L, as the 1000 ppm Standard Stock is used to perform standard addition.

- Place the electrode in the sample and press “next”. The ISE will start measuring, once stable it will show a prompt asking “stable?”, press “yes”. The next screen will prompt the analyst to “enter std. vol. in mL”, enter 0.5 mL for the volume of standard that is being added to the sample being run and press accept. The instrument will ask “std. added?”, spike the sample with the know volume of the stock standard stated previously, and press yes to begin the measurement.

- The ISE will be measuring the mV of standard + sample. Once stable, a prompt will appear asking “stable?”, press “yes”. The instrument will then show the concentration of F- in mg/L. Press “next”, if another sample needs to be run, press “yes” in the “repeat technique” screen that appears. If the analysis has concluded, press “no”.

- Calculations Modifications:

- Samples are not blank corrected. Follow management instructions for spectroscopy blank corrections.

Method Review

OSHA ID-110 Fluoride (F- and HF) in Workplace Atmospheres

- Date of Review:  August 17th by Hanh Dinh.

- Status of Method:  Fully Validated Method use is ongoing.  Approximately 187 samples were analyzed for fluoride (1280) and 141 samples were analyzed for HF (1460) in the last 5 years.

- Deviations from Method as Written: See the Method Work Instruction above.

- Formatting, grammar or spelling issues: Spell out “one hour” or “an hour” rather than the “1 h” as is printed in section 6.5.2 step (b) and section 6.5.3. Section 6.5.2 step (d): The reference to step (5) should be to step (e) instead. Section 6.7.2: Correct “Griff beakers” to “Griffin beakers”. Section 6.7.7 and 6.7.8 contradict each other on the standard addition spike, which should be a 500 µL aliquot of the 1000 µg/mL fluoride stock standard (equivalent to 500 µg F-).

Appendix A: Method Modification Validation Data

Table 1. Reporting Limit Analyst Spiked Media (ASM) Percent Recovery Results for IMIS 1460

| Sample Name | Solution Volume (mL) | Instrument Result (ppm F-) | Result (µg F-) | Spike Theoretical Value (µg F-) | % Recovery |
| --- | --- | --- | --- | --- | --- |
| RLV1 | 50 | 0.401 | 20.05 | 20.00 | 100.25 |
| RLV2 | 50 | 0.395 | 19.75 | 20.00 | 98.75 |
| RLV3 | 50 | 0.393 | 19.65 | 20.00 | 98.25 |
| RLV4 | 50 | 0.387 | 19.35 | 20.00 | 96.75 |

Table 2. QCSM-Level spiked samples Percent Recovery Results for IMIS 1460

| Sample Name | Solution Volume (mL) | Instrument Result (ppm F-) | Result (µg F-) | Spike Theoretical Value (µg F-) | % Recovery |
| --- | --- | --- | --- | --- | --- |
| 119916 | 50 | 1.76 | 88 | 90.56 | 97.17 |
| 119917 | 50 | 3.41 | 170.5 | 181.12 | 94.14 |
| 119918 | 50 | 6.7 | 335 | 362.24 | 92.48 |
| 118496 | 50 | 1.74 | 87 | 90.56 | 96.07 |
| 118497 | 50 | 3.34 | 167 | 181.12 | 92.20 |
| 118498 | 50 | 6.7 | 335 | 362.24 | 92.48 |
| 118500 | 50 | 1.73 | 86.5 | 90.56 | 95.52 |
| 118501 | 50 | 3.409 | 170.45 | 181.12 | 94.10 |
| 118502 | 50 | 6.65 | 332.5 | 362.24 | 91.74 |
| AQ1 | 50 | 2.09 | 104.5 | 113.1 | 92.40 |
| AQ2 | 50 | 4.13 | 206.5 | 226.2 | 91.29 |
| AQ3 | 50 | 8.14 | 407 | 452.4 | 89.96 |
| BQ1 | 50 | 2.05 | 102.5 | 113.1 | 90.63 |
| BQ2 | 50 | 4.1 | 205 | 226.2 | 90.63 |
| BQ3 | 50 | 8.32 | 416 | 452.4 | 91.95 |
| CQ1 | 50 | 2.07 | 103.5 | 113.1 | 91.51 |
| CQ2 | 50 | 4.1 | 205 | 226.2 | 90.63 |
| CQ3 | 50 | 8.38 | 419 | 452.4 | 92.62 |
| DQ1 | 50 | 2.08 | 104 | 113.1 | 91.95 |
| DQ2 | 50 | 4.08 | 204 | 226.2 | 90.19 |
| DQ3 | 50 | 8.32 | 416 | 452.4 | 91.95 |

*Each set of spiked QCs and RLVs was analyzed with a corresponding blank. All blanks were below the established Reporting Limit for the method and were not included in the above tables.